## Slide 2
Friederike EWALD
Beam Size Monitoring at the ESRF:
Comparison of different X-ray based techniques
Topical Workshop on Emittance Measurements for Light Sources and FELs
29th – 30th January 2018, Barcelona (ALBA)

## Slide 3: Contents
Current status of X-ray based beam size monitoring
  Simple projection of X-rays
  Pinhole Camera (imaging)
  X-ray Lenses (imaging)
Prospects and challenges for ESRF-EBS
  EBS beam parameters/implementation of DIAG X-ray beam ports
  Evaluation of beam size monitoring techniques for ESRF-EBS

## Slide 4: Existing Emittance Monitors at the ESRF
4
12
10
11
9
8
7
6
1
21
20
19
18
17
16
15
14
13
22
28
27
26
25
24
23
31
30
29
5
2
3
32
ID25
D9
12 X-ray Projection Monitors
3 Pinhole Cameras
X-ray Lens (XRL, removed)
D11
X-ray imaging
X-ray projection
X-ray sources = Bending magnet radiation
Pinhole, Lens
(ID30 – Undulator radiation)

## Slide 5: Dipole radiation X-ray Projection Monitors
Dipole
Electrons
X-rays
Cu-absorber
(35 mm)
Detector
Scintillator: Lu1.8Y.2SiO5:Ce (“PreLude”) 500 mm
Focusing via motorised scintillator screen
Double Achromat
CCD: Flea (Point Grey Research)
Pb - shielding
Simple measurement of the X-ray beam projected from the source point in a bending magnet onto a scintillator screen
EPEAK ~ 160 keV
s ‘ SYNC RAD = 17.5 mrad
Scintillator
Actuator for scintillator (focusing)
CCD & lenses in Pb shielding
Firewire control

## Slide 6: X-ray projection monitors – View from the top
Crotch absorber vessel
Vacuum pipe
X-rays
Vacuum pipe Electrons
Detector
Bending Magnet

## Slide 7: X-ray projection monitors – Synchrotron Radiation vs. Electron Beam Size
Total beam size on screen: S 2tot = eb - 2ea D + eg D2 + s’2ph D2)
electron beam size and divergence
divergence of synchrotron radiation
Chart: 
| 200 | 190 | 180 | 170 | 160 | 150 | 140 | 130 | 120 | 110 | 100 | 90 | 80 | 70 | 60 | 50 | 40 | 30 | 20 | 10 | 0 | -10 | -20 | -30 | -40 | -50 | -60 | -70 | -80 | -90 | -100 | -110 | -120 | -130 | -140 | -150 | -160 | -170 | -180 | -190 | -200
Total | 1.1139 | 11.9188 | 111.7381 | 918.0091 | 6616.8945 | 41929.1127 | 233471.6233 | 1144099.4523 | 4936647.3212 | 18767561.516 | 62882400.6102 | 185790719.8066 | 484489839.4055 | 1116855741.6489 | 2281210760.4206 | 4140589092.0445 | 6700756581.1347 | 9700994111.9335 | 12603337603.2635 | 14730438389.6384 | 15513541502.5939 | 14730438389.6384 | 12603337603.2635 | 9700994111.9335 | 6700756581.1347 | 4140589092.0445 | 2281210760.4206 | 1116855741.6489 | 484489839.4055 | 185790719.8066 | 62882400.6102 | 18767561.516 | 4936647.3212 | 1144099.4523 | 233471.6233 | 41929.1127 | 6616.8945 | 918.0091 | 111.7381 | 11.9188 | 1.1139
Electrons 5 pm | 5.915e-44 | 9.5562e-39 | 8.3468e-34 | 3.9415e-29 | 1.0063e-24 | 1.3889e-20 | 1.0364e-16 | 4.1811e-13 | 9.1192e-10 | 1.0753e-06 | 0.00068552 | 0.23627 | 44.0261 | 4435.202 | 241558.8539 | 7112769.5092 | 113229798.2257 | 974515955.2348 | 4534437260.2731 | 11406806205.5153 | 15513541502.5939 | 11406806205.5153 | 4534437260.2731 | 974515955.2348 | 113229798.2257 | 7112769.5092 | 241558.8539 | 4435.202 | 44.0261 | 0.23627 | 0.00068552 | 1.0753e-06 | 9.1192e-10 | 4.1811e-13 | 1.0364e-16 | 1.3889e-20 | 1.0063e-24 | 3.9415e-29 | 8.3468e-34 | 9.5562e-39 | 5.915e-44
Electrons 10 pm | 3.0292e-17 | 1.2176e-14 | 3.5984e-12 | 7.8196e-10 | 1.2494e-07 | 1.4679e-05 | 0.001268 | 0.080538 | 3.7613 | 129.1589 | 3261.118 | 60542.7246 | 826438.4508 | 8294919.5798 | 61216283.0034 | 332181042.4738 | 1325366052.8716 | 3888212149.1086 | 8387203385.4054 | 13302629870.8679 | 15513541502.5939 | 13302629870.8679 | 8387203385.4054 | 3888212149.1086 | 1325366052.8716 | 332181042.4738 | 61216283.0034 | 8294919.5798 | 826438.4508 | 60542.7246 | 3261.118 | 129.1589 | 3.7613 | 0.080538 | 0.001268 | 1.4679e-05 | 1.2494e-07 | 7.8196e-10 | 3.5984e-12 | 1.2176e-14 | 3.0292e-17
Electrons 30 pm | 19.3903 | 143.1001 | 953.1889 | 5730.6315 | 31096.4183 | 152300.9553 | 673254.486 | 2686211.8139 | 9673544.285 | 31442376.5183 | 92242162.1638 | 244246278.3892 | 583728376.8522 | 1259151617.9522 | 2451487143.2749 | 4307899693.3753 | 6832598723.1354 | 9781165759.9471 | 12638039044.8596 | 14738494954.5348 | 15513541502.5939 | 14738494954.5348 | 12638039044.8596 | 9781165759.9471 | 6832598723.1354 | 4307899693.3753 | 2451487143.2749 | 1259151617.9522 | 583728376.8522 | 244246278.3892 | 92242162.1638 | 31442376.5183 | 9673544.285 | 2686211.8139 | 673254.486 | 152300.9553 | 31096.4183 | 5730.6315 | 953.1889 | 143.1001 | 19.3903
Synchrotron radiation

## Slide 8: X-ray projection monitors – Comparison of data and model
<ez>meas = (20 ± 3) pm
< ez >mod = (20 ± 3) pm
      Alignment critical at low emittance
      Deconvolution obligatory
<ez>meas = (3 ± 1) pm
< ez >mod = (2 ± 0.3) pm
Detector Name (cell)

## Slide 9: Monitoring of 14 vertical emittance values during USM

## Slide 10: Average vertical emittance monitoring
Average vertical emittance serves for manual or automatic correction of the machine during operation.
Average <ez> very sensitive to small emittance changes (De ≈ 0.2 pm !)
Average ez
Lifetime (Libera)
Time
0.5 pm

## Slide 11: Contents
Current status of X-ray based beam size monitoring
  Simple projection of X-rays
  Pinhole Camera (imaging)
  X-ray Lenses (imaging)
Prospects and challenges for ESRF-EBS
  EBS beam parameters/implementation of DIAG X-ray beam ports
  Evaluation of beam size monitoring techniques for ESRF-EBS

## Slide 12: X-ray pinhole = Image formation through a small aperture
CCD
3 mm Al (watercooled)
Pinhole
V = 10um
H = 50um
Dipole
e-beam
Fluorescent screen
Lens
Air
Vacuum
d
D > d
Pinhole magnification M = D / d
S = M · s
Electron beam
profile
Beam profile on screen
s
Mirror

## Slide 13
Scintillator & CCD camera ~16 m downstream of pinhole
X-ray path in air
Al window
Pinhole
~ 4 m from source

## Slide 14: Pinhole area Detection area
Pinhole
Cu-attenuator
Al-window
N2 flushed pipes

## Slide 15: Pinhole Resolution
Pinhole
Point Source
Electron Beam profile
Screen
M = D / d
d
D
1. projection shadow (upper limit of pinhole size)
2. diffraction from pinhole (minimum pinhole size)

## Slide 16: Resolution of X-ray pinhole camera
Geometric shadow of the pinhole Diffraction from pinhole
Minimum between both effects leads to minimum resolvable source size and emittance :
@ D11:
smin = 6 mm with Aopt = 10 mm
ez, min = 1 pm
@ ID25:
smin = 5 mm with Aopt = 10 mm
ez, min = 0.5 pm

## Slide 17: X-ray Pinhole Screenshots

## Slide 18: Contents
Current status of X-ray based beam size monitoring using
  Simple projection of X-rays
  Pinhole Camera imaging
  X-ray Compound Refractive Lenses
Prospects and challenges for ESRF-EBS
  EBS beam parameters/implementation of DIAG X-ray beam ports
  Evaluation of beam size monitoring techniques for ESRF-EBS

## Slide 19: X-ray Lens = Image formation with stack of parabolic concave lenses
Focal length of CRLs (thin lens approximation):
R… Radius of curvature
N… Number of lenses
d… Index of refraction (real part)
E… Photon energy
Lens material: Aluminium
nAl << nair for X-rays
Image from [1]
2 mm
the lens

## Slide 20: X-ray Lens = Image formation with stack of parabolic concave lenses
3 mm Al (water cooled)
CRL
Dipole
e-beam
Air
Vacuum
L1
L2
CRL magnification M = L2 / L1
S = M s
Electron beam
profile
Beam profile on screen
s
CCD
monochromator

## Slide 21: Emittance beamline D11 equipped with CRLs
Si(111) monochromator
(Laue configuration)
Scintillator,
Lenses,
CCD
Pinhole
CRL

## Slide 22: Choosing the number of lenses
focal length (thin lens approximation):
f = R/2Nd N... number of single lenses
d... index of refraction real part
32 keV
 for example 13 lenses at Eopt = 32 keV (used at emittance beamport D11)

## Slide 23: XRL for emittance measurement at ESRF ( now removed )
[1] B. Lengeler et al. / Nuclear Instr. Meth. in Phys Res. A 467-468 (2001) 944-950
[2] B. Lengeler et al. /J. of Appl. Phys. 84 (1998) 5855
[3] http://purple.ipmt-hpm.ac.ru/xcalc/xcalc_mysql/crl_par.php
Transverse resolution determined by diffraction on lens aperture :
smin = 0.75 · L1 /Deff [1]
Deff ...... effective aperture of the lens determined by geometric aperture, absorption
and surface/shape imperfections.
Diffraction limited resolution of lens at beam port D11 ( calculated using [2] or [3] ):
    smin = 3.3 mm
    ez,min = 0.3 pm

## Slide 24: Comparison Pinhole – CRL (at beamport ID25)

## Slide 25: Contents
Current status of X-ray based beam size monitoring using
  Simple projection of X-rays
  Pinhole Camera imaging
  X-ray Compound Refractive Lenses (CRL)
Prospects and challenges for ESRF-EBS
  EBS beam parameters/implementation of DIAG X-ray beam ports
  Evaluation of beam size monitoring techniques for ESRF-EBS

## Slide 26: Challenges and prospects for emittance measurement for ESRF-EBS
ESRF – EBS
ESRF
2020
ex = 4000 pm
ex = 130 pm
30
2019
Horizontal emittance
Vertical emittance
ez < 1 pm !!!

## Slide 27: ESRF-EBS
  Wish list for the perfect emittance beamline/source point location:
  pinhole close to source = high magnification factor
  strong B-field = high photon flux
  high beta-function
  ideally 2 locations with different dispersion for energy spread measurement

## Slide 28: EBS emittance beam port positions:
Source 2: DQ1D
(3x)
Pinhole position 2:
close to diagnostics drift section DR37
5.4 m
d = 3.76 m
28
D = 16.35 m
Screen
M = 4.35
Source 1: DL1A_5
(2x)
Pinhole/lens position 1: close to BPM03

## Slide 29: Mechanical Assembly for Emittance ID beam port
X-rays
Electrons

## Slide 30: Emittance – Challenge: small beam sizes
 | Expected for ESRF EBS |  |  |  | Pinhole | CRL
 | sx [mm] | sz [mm] |  |  | smin [mm] | 
 | (ex = 130 pm) | coupling 4% (ez = 5 pm) | coupling 1% (ez = 1.3 pm ) | coupling 0.1% (ez = 0.13 pm ) |  | 
 |  | USM conditions |  | Perfect machine |  | 
DL1A_5 bz = 16.64 m | 15 | 9 | 4.7 | 1.5 | 5 | 3
DQ1D bz = 3.81 m | 19 | 4 | 2.2 | 0.7 | 6 | 2.5
Small emittance + Small beta-function  Small source beam size
Horizontal emittance : 130 pm

## Slide 31: X-ray Fresnel diffraction from a slit
Simulations with SRW for the ESRF low emittance lattice:
Optimum slit size for obtaining a
« simple » double-lobed diffraction pattern [1]:
d = 3.76 m D = 16.35 m
Ebeam = 6 GeV A = 200 mA
Eph = 40 keV ex = 0.11 pm
a = 26 mm
Dipole
Slit
Monochromator
Observation
plane
[1] M. Masaki et al., “X-ray Fresnel diffractometry for ultralow emittance diagnostics of next generation synchrotron light sources”, PR Special Topics – Accelerators and Beams, 18, 042802 (2015)
d
D

## Slide 32: X-ray Fresnel diffraction from a slit
Simulations with SRW for the ESRF low emittance lattice:
Chart: Vertical position on screen [m]
a = 26 um
Chart: 
| 0 | 0.1 | 0.2 | 0.30000000000000032 | 0.4 | 0.5 | 0.60000000000000064 | 0.70000000000000062 | 0.8 | 0.9 | 1 | 2 | 3 | 4 | 5 | 6 | 7
Visibility | 0.51387 | 0.39141 | 0.37148 | 0.35286 | 0.33529 | 0.31876 | 0.30351 | 0.28942 | 0.27617 | 0.26403 | 0.25273 | 0.17207 | 0.12168 | 0.085925 | 0.059544 | 0.039805 | 0.025139

## Slide 33: Summary
Simple X-ray projection monitors can be a nice way of implementing a multitude of small, cheap and robust emittance monitors for the vertical plane, but with limited resolution.
The X-ray pinhole is a very robust method for emittance measurement with a very good resolution.
However, for the ultra small vertical emittances expected for ESRF-EBS its resolution is not sufficient.
X-ray compound refractive lenses have a slightly improved theoretical resolution limit, but even they will not able to resolve electron beam sizes of < 3 um.
X-ray Interference techniques may be the solution to measure the smallest expected vertical emittance of < 1 pm.

## Slide 34: Not adressed here:
Damage induced by ozone production (X-rays + O2  O3  corrosion)
Degradation of CCD /CMOS cameras due to radiation  Test and choice of cameras
Very small source sizes mean also very small image sizes  VIS-imaging has to be designed in consequence.
Vibration (electron beam and beamline elements) will start to be a problem at ultra small beam sizes.

## Slide 35
Thank you for your attention !
For more information on the ESRF-EBS upgrade:
ESRF website
Technical design report “Orange book”:
http://www.esrf.fr/home/about/upgrade/upgrade-programme-documentation/orange-book.html
Articles on selected topics on the ESRF-EBS blog : http://ebs.esrf.fr/