## Slide 1: PLANNED X-RAY DIFFRACTION DIAGNOSTICS FOR APS-U EMITTANCE MEASUREMENTS
Source
Bingxin Yang and Soonhong Lee
Advanced Photon Source
Argonne National Laboratory
Emittance Measurements for Light Sources and FELs
January 29-30, 2018
Notes: Univ of Chicago Review, Aug. 30, 2010
1

## Slide 2: Outline
Emittance monitor specifications
Right x-ray source is half the success
Design of x-ray diffraction apertures
Simulation of diffraction profiles and estimation of beam size monitor performance
Key features of optical / mechanical design
Summary

## Slide 3: APS-U Storage Ring Emittance Monitor
Operating Mode | Horizontal Emittance | Vertical Emittance | Beam Current
Timing Mode | 31.9 pm | 31.7 pm | 200 mA
Brightness Mode | 44 pm | 4.2 pm | 200 mA
Initial Commissioning Mode |  |  | 25 mA
Beam Instability? | 1000 pm | 500 pm | 
Low emittance ratio study (1%)? |  | 0.4 pm | 
Expected storage ring emittance*
Horizontal Emittance | Vertical Emittance | Beam Current
25 – 1000 pm | 0.4 – 500 pm | 20 – 200 mA
Emittance monitor operating requirements*
* Two planned operating modes for user operation
* Planned bandwidths: 0.1 – 1 Hz for absolute beam size monitor; 1 – 10 Hz for relative beam size monitor.

## Slide 4: X-ray Source for Beam Size Monitor
SR emittance will be calculated from measured beam sizes. Choosing proper x-ray source is a critical step to the successful emittance monitor! The APS-U storage ring chamber has two light ports, looking at 7 – 8 x-ray source points:
ID port: insertion devices + bend magnet sources B:M1 and A:M1
BM port: M4, Q8 (too weak), M3, and optional three-pole wiggler
APS-U storage ring lattice functions.
ID
A:M1
B:M1
A:M4
3PW
B:M3

## Slide 5: X-ray Source for the Emittance Monitor
Port | ID |  | BM |  | 
Source | A:M1/B:M1 | ID | M4 | 3PW | M3
X-size from emittance | 7.2 | 12.5 | 5.0 | 6.4 | 4.7
X-size from energy-spread** | 0.4 | 0.9 | 8.0 | 11.0 | 4.9
Horizontal beam size (X) | 7.2 | 12.6 | 9.4 | 12.8 | 6.8
Vertical beam size (Y)* | 9.1 | 2.8 | 4.3 | 3.2 | 4.5
Minimum source sizes for APS-U user operations*
[Observations]
At the BM port, the source sizes are dominated by contributions from beam energy spread, making pure emittance measurement difficult!
At the ID port, the vertical ID source size is small, difficult to measure!
The B:M1 is 7.7 m upstream of A:M1, farther away from the x-ray optics!
The A:M1 source is ideal for emittance measurement!
* Minimum emittances in user operations are 31.9pm / 4.2pm in horizontal / vertical planes.
** Nominal electron beam energy spread is 0.136%.

## Slide 6: APS-U Storage Ring Beam Size Monitor
Operating Mode | Horizontal Beam Size | Vertical Beam Size | Beam Current
Timing Mode (User Op) | 7.3 mm | 25.1 mm | 200 mA
Brightness Mode (User Op) | 8.5 mm | 9.1 mm | 200 mA
Initial Commissioning Mode |  |  | 25 mA
Beam Instability? | 41 mm? | 100 mm | 
Low emittance ratio study (1%?) |  | 3 mm | 
Expected electron beam sizes at A:M1 source point*
Horizontal Beam Size | Vertical Beam Size | Beam Current
6.5 – 41+ mm | 3 – 100+ mm | 20 – 200 mA
Beam size monitor operating range requirements*
* Two planned modes for user operation
* Planned bandwidths: 0.1 – 1 Hz for absolute beam size monitor; 1 – 10 Hz for relative beam size monitor.

## Slide 7: X-ray Diffraction Optics
The x-ray optics is a 0.1-mm tungsten foil with six apertures:
  Aperture spacing is 0.5 mrad (3.3 mm), making separate images at detector
  The apertures are located 6.56 m from the source
  Apertures may be put into two function groups.
Absolute beam size monitor (ABSM): obtain beam sizes by analyzing x-ray diffraction profiles, operating in 0.1 – 1 Hz range. Three apertures cover different beam size ranges:
  Monochromatic x-ray pinhole camera (15 keV): for 8 – 100+ mm beam sizes.
  Wide-aperture Fresnel diffractometer (8 keV): 4 – 14 mm
  Young’s double slits interferometer (8 keV): 2 – 6 mm
Relative beam size monitor (RBSM): obtain beam size information by monitoring x-ray diffraction peak intensities, operating at 1 – 10+ Hz:
  Double-slits collimator for horizontal beam size (15 keV): 4 – 100 mm
  Double-slits collimator for vertical beam size (15 keV): 4 – 100 mm
  X-ray beam position monitor (15 keV) for maintaining collimator alignment

## Slide 8: 15-keV X-ray Pinhole Camera: Geometry
Today, most x-ray pinhole-camera emittance monitors use white beams. But the monochromatic beam has several unique advantages:
Diffraction profiles can be more precisely modelled, resulting in more accurate beam size data.
Photon energy can be chosen in a opaque region for tungsten, making the pinhole slits easier to make, and with a better optical quality.
Pinhole camera profiles for source size 2 – 18 mm (left panel) and 5 – 80 mm (right panel)
Source distance | Pinhole | Image distance | M | Source sizes
6.6 m | 34 mm × 34 mm | 13.4 m | 2.6 | 8 – 100 mm

## Slide 9: 15-keV X-ray Pinhole Camera: Data Analysis
For small sources, full diffraction profiles get best curve-fitting results.
Shortcut for larger beam sizes (> 8 mm): Fit 60% of the peak data to a Gaussian curve. The resulting peak width sG can be “corrected for resolution effect” using the following expression:
with
The resolution-corrected source sizes carry an error below 1 mm for source size < 8 mm
X-ray absorption by tungsten is strong at 15 keV. A 50-mm foil has less than 0.01% transmission.
0.001
-0.002

## Slide 10: 8-keV X-ray Fresnel Diffractometer
If we widen the diffraction aperture to include the second Fresnel zones, the diffraction profile will have a center minimum. The entire profiles can be used for image analyses to extract beam size information.
Fresnel diffractometer profiles for source size 2 – 18 mm
Source distance | Pinhole | Image distance | M | Source sizes
6.6 m | 70 mm × 70 mm | 13.4 m | 2.6 | 4 – 14 mm

## Slide 11: 8-keV Young’s Interferometer (Vertical)
For very small beam sizes, a pair of double slits will be used to make a Young’s Interferometer, which produces sinusoidal interference fringes.
Full diffraction profiles will be used for image analyses.
Expect longer integration time
If necessary, a larger slit spacing can be used for even smaller beam sizes
Young’s interference pattern from two 8-mm slits with 80-mm spacing
Source distance | Slits | Image distance | M | Y-Beam sizes
6.6 m | 2 × 8 mm 80 mm spacing | 13.4 m | 2.6 | 2 – 6 mm (0.2 – 1.8 pm)

## Slide 12: Relative BSM: Double-Slits Collimator
One-dimensional 15-keV x-ray pinhole camera:
Pinhole-slit width is chosen to maximize the peak intensity at the detector.
The slits’ length increases the x-ray flux by five fold (relative to a pinhole).
Detector slits width is chosen to balance good resolution and good signal level.
Concept of the APS-U relative beam size monitor
Source distance | Pinhole slits | Detector distance | Detector slits | M | Source sizes
6.6 m | 34 mm × 150 mm | 13.4 m | 10 mm × 400 mm | 2.6 | 4 – 100 mm

## Slide 13: Relative BSM: Two Modes of Operations
Scanning mode: scanning a narrow slit across the pinhole image
  A slow, but accurate absolute beam size monitor, free from degradation by x-ray imaging detectors
  Beam size obtained by fitting the entire intensity profile  Reliable calibration
Intensity mode: measure area and peak signal intensities
  In user operations, the area-to-peak intensity ratio will be used to monitor beam size /emittance variations for emittance feedback system.
Relative Beam Size Monitor detector slits scanning profiles and peak intensity depends on source sizes.

## Slide 14: RBSM: Alignment Requirements
In the vertical plane, the pinhole slits needs to be centered in the BM x-ray fan.
In both horizontal and vertical planes, the two-slit collimators need to be aimed precisely at the source. Any misalignment will cause detector signal drop, which may be misinterpreted as source size increase.
To keep detector centered within top 2% of the peak intensity, we need to center the collimator within 1 mm from the source center.
Relative Beam Size Monitor detector slits scanning profiles indicate alignment requirements.

## Slide 15: X-ray Beam Position Monitors
XBPM1 (for beam angle) is mounted on the x-ray aperture holder.
XBPM2 (for source point) shares the same mono and support with the RBSM detectors.
XBPM1 Ta blades
X-ray apertures
Detector Slits
BEAM
Si(111) crystal
Si-PIN photodiode
XBPM1 is for centering the bend magnet fan vertically on the aperture array
  Photoemission XBPM design integrated with the aperture holder (@6.56 m)
  Resolution of several microns is more than adequate for BM fan height of 500 microns.
XBPM2 is for aligning the source, pinhole slits, and the detector slits on the same straight line.
  Large pinhole (300 mm × 200 mm) between XBPM1 blades
  15-keV mono and 2-D position-sensitive detector to readout source point position
  Centering beam in the XBPMs centers the beam in the relative beam size monitors!

## Slide 16: Implementation: Optics & Mechanics
Beam Size Monitor will use S38A:M1 as the source, sharing the same straight section with the RF cavities in S37ID.
Enclosure’s radiation shielding managed 37-ID Front End.
RF Cavity (Z = 0 m)
A:M1 source (3.84 m)
X-ray apertures (10.40 m)
Photon shutter and valve
Windows, monochromators, detector and cameras (18.5 – 20 m)
Radiation enclosure equivalent to 37-ID front end

## Slide 17: Mechanical Design: Vibration Control
“Vibration-free” mounting for the pinhole apertures:
Water-cooled absorber incepts ~75% beam power, lowers aperture flow req.
Water-cooled x-ray apertures are mounted rigidly on the storage ring chamber.
Reduce chamber flexing by placing support column under the apertures.
Estimated RMS vibration amplitude < 0.1 mm. XBPM handles slow motion.
X-ray Beam
Old Support Column
Power Absorber (>75%)
New Support Column
Water-Cooled Pinhole Apertures (6.56 m from source)
Water-Cooled GlidCop absorber

## Slide 18: X-ray Optics Design
Coherence preservation is the most important consideration:
Place transmission optics (window, lens, filters,…) at the very end.
Place reflective optics (mirrors, monochromators,…) at the very end.
X-ray photons travel in vacuum in the first 92% of the beamline. Its wavefront only disturbed by diffraction apertures. Aperture edge perfection is important.
To preserve coherence, the first window is placed 1.5 M upstream of the last YAG screen
Be Window
I0 Monitor
Detector slits
Si(111) Mono
Si-PIN PD
15-keV Mono
8-keV Mono
Mono Be Window
YAG-1
YAG-2
sCMOS-1
sCMOS-2
1.5 M

## Slide 19: Summary
The A:M1 source is very important to successful APS-U storage ring emittance diagnostics:
  Low dispersion allows for clean emittance measurements
  Larger beam sizes relax resolution requirements (>2x)
An array of six x-ray diffraction apertures will be used:
  For absolute beam size measurements, we will use a pinhole camera, a wide aperture Fresnel diffractometer and a Young’s double-slit interferometer.
  For relative beam size changes, 1-D double-slit collimator will be used to monitor normalized peak intensities.
Coherence preservation is the most important concern.
The diagnostics cover all expected storage ring beam conditions, including extremely small vertical beam emittance (< 0.4 pm). Upgrade to 0.1-pm is possible with different set of double slits.