## Slide 1: VADER: Variable Aperture Dipole for Elettra Ring
M. Domínguez, F. Toral (CIEMAT), Y. Papaphilippou, Axel Poyet (CERN), Emanuel Karantzoulis (Elettra), Raffaella Geometrante (Kyma)

## Slide 2: Summary
Where we started: CLIC Prototype
What we were asked for: VADER Specs
What we did: New Design
Where we are: Current Status
What we expect: Next steps

## Slide 3: Where we started: CLIC Prototype
SmCo in LF
NdFeB in MF and HF
Flux concentration
Hyperbolic pole tips profile
Armco
Fe-Co (Vacoflux)
Flat pole tip profile
Field trimming
Dimensions:
0.65 x 0.68 x 0.56 m
Weight: 1.2 T
Steel 1010
Combined Function:
Longitudinal gradient with trapezoidal shape (2.3 T Peak)
Transverse gradient 11 T/m
CLIC DRs Main Bending Magnet (Prototype)

## Slide 4: Where we started: CLIC Prototype
Notes: Once the preliminary assembly was done and everything was checked, the final assembly, with the permanent magnets, started

## Slide 5: Where we started: CLIC Prototype
Proof of concept: First accelerator magnet with longitudinal trapezoidal gradient! Key factor in terms of beam emittance reduction
Very good field quality achieved
Quadrupolar component achieved
Trapezoidal longitudinal field achieved… …With an issue
Explaining theory: missing anneal in the LF and MF poles (Armco)
Two new annealed poles produced and installed in one of the LF modules
Test bench results showed exactly the same issue as before
The permanent magnet volume is increased around 30%,reaching the demanded field peak of 2.3T with a magnet gap of 19mm (17+2mm)
Chart: 

## Slide 6: What we were asked for: VADER Specs
Task 7.3 within I.FAST WP7: High Brightness Accelerators for Light Sources
  Funds from Horizon 2020 Innovation Pilot Project for Particle Accelerators, ranked 3rd in the Prototypes proposals category
VADER joint project was born:
 | CLIC Prototype Specs | VADER (Variable Aperture Dipole for Elettra Ring) Specs
Magnet Gap (mm) | 13 | 19
Max Field (T) | 2.3 | 2.3
Transverse G (T/m) | 11 | 23
Field Quality (Units 1E-4) | ≈ 1 | ≈ 1
Magnet Length (cm) | 56 | 80
Magnet Specs
Magnetic & Mechanical Design
Production & Assembly
VADER optimal field profiles

## Slide 7: What we did: New Design
Minor changes in the philosophy, major changes in the poles (90% of the design time)
All NdFeB permanent magnet blocks, including LF modules
The permanent magnet volume is increased around 40%, reaching the demanded field peak of 2.3T with a magnet gap of 19mm (17+2mm)
New field trimming design implemented
    Yoke completely split in two parts supported by an aluminium block
    Sliding parts perform the regulation
    Higher/better regulation than the CLIC prototype
CLIC poles
VADER poles

## Slide 8: What we did: New Design
VADER Overview
NdFeB in LF, MF and HF
Flux concentration
Hyperbolic pole tips profile
Armco
Fe-Co (Vacoflux)
Flat pole tip profile
Field trimming
Dimensions:
0.65 x 0.68 x 0.80 m
Weight: 1.5 T
Steel 1010
Combined Function
Longitudinal gradient with trapezoidal shape (2.3 T Peak)
Transverse gradient 23 T/m

## Slide 9: Where we are: Current Status
The field quality is still very good. Comparable to the CLIC prototype
Longitudinal gradient profile achieved
Transverse gradient difficult to reach, the present design is 19 T/m
Meeting with CERN pending to check the viability of the present design

## Slide 10: What we expect: Next steps
Mechanical Design

## Slide 11: What we expect: Next steps
Mechanical Design:

## Slide 12
Thank you for your attention!