## Slide 1
CHALLENGES WITH THE PERMANENT MAGNETS DIPOLES DEVELOPMENT FOR ESRF-EBS
Chamseddine BENABDERRAHMANE
On behalf of Insertion Devices and Magnets Group
Workshop on Permanent Magnets for Low Emittance Rings
Trieste, Italy, 14th and 15th November 2023
bchams@esrf.fr

## Slide 2: OUTLINE
Introduction
Motivation
Design and Prototype
Procurement of PM and mechanical parts
Dipole assembly and magnetic measurements
Commissioning and Operation
Conclusion

## Slide 3: Introduction
Magnet layout in one cell
Combined dipole-quadrupole
Sextupole
Octupole
Moderate gradient quadrupole
High gradient quadrupole
Permanent magnet dipole
with longitudinal gradient (DLs)

## Slide 4: OUTLINE
Introduction
Motivation
Design and Prototype
Procurement of PM and mechanical parts
Dipole assembly and magnetic measurements
Commissioning and Operation
Conclusion

## Slide 5: Motivation
Compact devices
No power supply and no cooling systems
Better reliability (no water and power supply failures)
Less control systems, cables and noise
Important reduction in operation cost
PM dipole have advantages over electromagnet one
Challenges for permanent magnet Dipoles
Magnetic field design
Magnetic field tuning and shimming
Temperature dependence
Demagnetization risks
Series production

## Slide 6: Motivation
Demonstrates feasibility of low cost permanent magnet dipole production
        Low procurement cost
        Energy efficient, no running cost
Full control on the dipole assemblies
In house experience in permanent magnet assemblies
In house experience in magnetic measurement
Magnetic tuning
Add a new experience in terms of large series of magnet assemblies and tuning

## Slide 7: OUTLINE
Introduction
Motivation
Design and Prototype
Procurement of PM and mechanical parts
Dipole assembly and magnetic measurements
Commissioning and Operation
Conclusion

## Slide 8: Design and prototype
Dipole |  | 
Gap | mm | 25.5 – 30.5
Iron length | mm | 1788
Permanent magnet |  | Sm2Co17
Iron |  | Pure iron
Number of dipoles |  | 128
 |  | DL1 | DL2
Strength | T | 0.67 – 0.17 | 0.54 – 0.17
Vertical field vs. longitudinal position
244 mm
183 mm
Courtesy of J. Chavanne

## Slide 9: Design and prototype
Motivation
DL1 | M5 | M4 | M3 | M2 | M1
Top magnet | 16 | 0 | 0 | 0 | 8
Side magnet | 20 | 20 | 16 | 12 | 0
DL2 | M5 | M4 | M3 | M2 | M1
Top magnet | 10 | 0 | 0 | 0 | 8
Side magnet | 20 | 20 | 16 | 12 | 0
Top magnet
Side magnet
Pole for M1
Yoke
M1
M2
M3
M4
M5
Pole for M2 to M5
Quarter Top magnet for field adjustment

## Slide 10: Design and prototype
Permanent magnet assembly
Magnetic field strength and quality
Longitudinal field integral fringe
Temperature compensation
Two modules with 0.62 T and 0.41 T
Prototypes to confirm calculated performance and
to define the series production process

## Slide 11: Design and prototype
PM assembly needs special tools
Easy and fast mechanical correction
Shimming required to reach targeted field
Longitudinal gap to be defined for flat field
Tolerance: DB/B < 10-3 @13 mm
gs
Flat field at longitudinal gap gs = 4 mm
Module 1 with 20 µm shims (Hall probe meas.)

## Slide 12: Design and prototype
Dominated by PM material temperature coefficient
PM devices with fixed field should be compensated
Compensated by passive Fe-Ni shunts
dB/B < 10-5/C after compensation
Field integral measurements on PM DL modules NdFeB PM, Sm2C017 PM
dB/B/dT after compensation:< 40 ppm/C
Pole
Yoke
PM block
Special Fe-Ni shunt
(low Curie temperature)
Material | dB/dT
Sm2Co17 | 3.3 10-4
Nd2Fe14B | 10-3
THERMOFLUX (VACUUMSCHMELZE)

## Slide 13: Design and prototype
Four DL modules tested
  Nominal field integral
  Temperature stability
  Br=1.067 T
Nominal field integral
  Excellent agreement between measurements and simulations
  Measured values are 0.2 ~ 1 % above targeted values
  Easy field integral reduction using soft iron shunts
Temperature stability after compensation:< 40 ppm/C
Prototype module

## Slide 14: OUTLINE
Introduction
Motivation
Design and Prototype
Procurement of PM and mechanical parts
Dipole assembly and magnetic measurements
Commissioning and Operation
Conclusion

## Slide 15: Procurement of PM and mechanical parts
Page 15
Permanent magnet procurement
Sm2Co17 permanent magnet
Around 6000 kg of permanent magnet (15130 blocks)
Produced by Magsound (China)
Sorting of magnets before delivery
Mechanical parts procurement
660 assembled empty modules to be delivered to ESRF
3 CFTs for the mechanical parts
(Raw material, Cut and annealing, Machining)
440 modules by AMF (UK) and 220 by Cecom (Italy)
Machining and assembly
Low carbon steel raw material
Flame cut

## Slide 16: OUTLINE
Introduction
Motivation
Design and Prototype
Procurement of PM and mechanical parts
Dipole assembly and magnetic measurements
Commissioning and Operation
Conclusion

## Slide 17: Dipole assembly and magnetic measurements
Page 17
Chartreuse Hall
Implementation in Chartreuse Hall
Dedicate area for dipole assembly (300 m2)

## Slide 18: Dipole assembly and magnetic measurements
Page 18
Dipole assembly area
1
2
3
4

## Slide 19: Dipole assembly and magnetic measurements
Page 19
Page 19
Modules assembly
Installation of the module on the assembly tool
Assembly of the permanent magnets in the module
Alignment of the poles
Installation of the thermal shims
Storage of the modules in the assembly area
Courtesy of F. Villar
Easy to use tooling
Management of magnetic forces
Robustness for long term use
One assembly tool for each dipole type

## Slide 20: Dipole assembly and magnetic measurements
Page 20
Modules magnetic measurement and shimming
Installation of the module on the magnetic bench and alignment
Two type of shimming are used, iron shunt for M2, M3 and M4, End iron shunt for M1 and M5
Measurement of the field integral and roll angle
Calculation and installation of the magnetic shims to reach the targeted field integral
Measurement for shims efficiency control, new shimming iteration if needed
Iron shunt

## Slide 21: Dipole assembly and magnetic measurements
Page 21
Vertical Integral Target : -253.5954 T.mm
Vertical Integral : -253.6065 T.mm
Horizontal Integral : 0.0634 T.mm
DB/B | Tolerance | Measured
-13 mm | -0.0025 | -0.0024719
13 mm | -0.0025 | -0.0016835
Module: DL1_5_18

## Slide 22: Dipole assembly and magnetic measurements
Page 22
Installation of the module on the top plate after the magnetic measurements using an assembly tool
The position of each module on the top plate is done with 3 pins
DL assembly

## Slide 23: Dipole assembly and magnetic measurements
Page 23
Page 23
DL magnetic measurement and shimming
Installation of the DL on the magnetic bench and alignment
Only End iron shunt (shims) on M1 and M5 are used
Measurement of field integrals and roll angle
Calculation and adjustment of the magnetic shims
Control the shims efficiency.
End iron shunt

## Slide 24: Dipole assembly and magnetic measurements
Page 24
Vertical Integral Target : -637.3030 T.mm
Vertical Integral : -637.2902 T.mm (2.10-5 of full DL)
Horizontal Integral : -0.0472 T.mm
Roll angle : -0.072076 mRad
Horizontal double Integral Target : -0.56324 T.mm2
Horizontal double Integral : -0.56326 T.mm2
Vertical double Integral : -5.018 10-5 T.mm2
Dipole: DL1E_13_12
DB/B | Tolerance | Measured
-13 mm | -0.0012 | -0.00051
13 mm | -0.0012 | -0.00087

## Slide 25: Dipole assembly and magnetic measurements
Page 25

## Slide 26: Dipole assembly and magnetic measurements
Page 26
Point | X mes | Y mes | Z mes
 | (mm) | (mm) | (mm)
module1 E | -815,27 | -134,51 | 358,61
module1 S | -615,21 | -131,42 | 358,55
module2 E | -457,77 | -129,79 | 358,42
module2 S | -257,76 | -128,54 | 358,43
module3 E | -100,01 | -128,09 | 358,48
module3 S | 100,01 | -128,14 | 358,49
module4 E | 257,71 | -128,56 | 358,48
module4 S | 457,73 | -129,52 | 358,46
module5 E | 615,42 | -130,55 | 358,46
module5 S | 815,43 | -132,21 | 358,49
Fiducialisation is done on each dipole using a Laser tracker
DL fiducialisation

## Slide 27: Dipole assembly and magnetic measurements
Page 27
DL assembly on mechanical support and packaging
Installation of the DL on the mechanical support
Scan and storage of all the quality documents
Plastic protection on the modules
Packaging in the wooden box
Storage in the Hall Chartreuse and ID14

## Slide 28: OUTLINE
Introduction
Motivation
Design and Prototype
Procurement of PM and mechanical parts
Dipole assembly and magnetic measurements
Commissioning and Operation
Conclusion

## Slide 29: Commissioning and operation
Page 29
Large discrepancies between model and real lattice at the restart
Errors in the calibration of the quadrupole magnets
Cross talk between magnets generation quadrupole errors
Quadrupole errors modeled by a thin lens with opposite polarity
No change of the lattice performance at the end
Calibration coefficients
for one magnet family

## Slide 30: Commissioning and operation
Page 30
PM DL
Quadrupole yoke
Quadrupole yoke
Sextupole yoke
Quadrupole yoke
Octupole yoke
4.7 cm
7.5 cm
6 cm
Short distance between magnet yokes
Less than 1 cm between sextupole and quadrupole coils
and between octupole and quadrupole coils
Saturated magnets
All ingredient for a cross-talk between magnets

## Slide 31: Commissioning and operation
Page 31
Dipole to quadrupole cross talk done in 2017
Needed for PM dipole tuning
Reduction in the integrated field of DL by 1%
Quadrupole to dipole cross talk not investigated
Dipole to quadrupole Cross-talk
Quadrupole to Dipole cross talk done 2020
0.9 % reduction in the integrated gradient
Effect localized at the dipole edge
Almost no change in the quadrupole
Quadrupole to dipole Cross-talk
(G Le Bec, LER2020, Frascati, 2020)

## Slide 32: Commissioning and operation
Page 32
RF frequency shift is following seasonal almost-reproducible variations.
Horizontal steerers strengths not drifting, average is kept to zero to use RF frequency for SR length variations
Winter SD
Winter SD
Summer SD
Summer SD
Imagine the SR as a VERY SLOW BOOSTER RAMPING DOWN in ENERGY. RF will not change. Beam ENERGY would reduce to adapt to the lower dipole fields in the SR.
We may then look at the AVERAGE QUADRUPOLE correction strengths (sextupoles are modified often for lifetime optimization and chromaticity setting).
After 3 years of operation it is not possible to observe clearly any loss of field in the permanent dipoles.
Courtesy of S.M Liuzzo

## Slide 33: Commissioning and operation
Page 33
DLs in operation since almost 3 years without any intervention

## Slide 34: OUTLINE
Introduction
Motivation
Design and Prototype
Procurement of PM and mechanical parts
Dipole assembly and magnetic measurements
Commissioning and Operation
Conclusion

## Slide 35: Conclusion
Simple magnetic structure, no electrical power and no cooling
Prototypes to solve all technical issues (assembly, measurement, tuning)
Procurement of all the parts including row material
Development of methods and technics for a large scale production
Quasi-industrial organization for the series production (workflow, procedures, quality insurances,…..)
Assembly of all the DLs in less than one year
3 years of operation without any intervention
Page 35

## Slide 36: Many thanks for your attention
Page 36