## Slide 1
Button Pickups for High
Bandwidth Applications
B. E. J. Scheible1,3, S. Vilcins2, A. Penirschke1
This work is supported by the German Federal Ministry of Education and Research (BMBF) under contract no. 05K22RO2.
1Technische Hochschule Mittelhessen, Wilhelm-Leuschner-Str. 13, 61169 Friedberg
2Deutsches Elektronen-Synchrotron DESY, Notkestr. 85, 22607 Hamburg
3Technische Universität Darmstadt, Schlossgartenstr. 8, 64289 Darmstadt

## Slide 2: Acknowledgements
There are many contributors who pathed the way:
The project partners, e.g., M.K. Czwalinna, N. Nazer, H. Schlarb [DESY]; E. Bründermann, A. Kuzmin, A.-S. Müller, G. Niehues, A. Schwarzenberger [KIT]; W. Ackermann, H. De Gersem [TU Darmstadt] and B.E.J. Scheible, A. Penirschke [THM];
S. Vilcins, who gave a lot of valuable input for designing a vacuum-sealed demonstrator;
Measurements at HZDR, M. Freitag, M. Kuntzsch, U. Lehnert, K. Zenker (+many more);
In the prior projects, leading to the cone-shaped pickups: A. Angelovski, V. Arsov, S. Bou Habib, M. Bousonville, M.K. Czwalinna, H. Dinter, M. Gensch, Ch. Gerth, M. Hansli, S. Hunziker, R. Jakoby, S. Korolczuk, A. Kuhl, M. Kuntzsch, A. Penirschke, K. Przygoda, J. Rönsch-Schulenburg, J. Roßbach, R. Rybaniec, H. Schlarb, V. Schlott, S. Schnepp, B. Steffen, C. Sydlo, J. Szewinski, S. Vilcins, M. Viti, T. Weiland (+many more);
We are grateful that Keysight made their Oscilloscope UXR1102A available to us & provided technical support through Thomas Kirchner, and that Neue Technologien GmbH & Co. KG (NTG) offered close cooperation during construction of the prototype even giving us the opportunity to carry out the final assembly on their premises.
This work is supported by the German Federal Ministry of Education and Research (BMBF) under contract no. 05K22RO2.

## Slide 3: Outline
Electro-Optical Bunch Arrival-Time Measurement
Pickups for Bunch Arrival-Time Monitors
  From button-pickups to 40-GHz cone-shaped pickups
Pickups on a Printed Circuit Board (PCB)
  Concept for a PCB-BAM
  Production of a first demonstrator
  Measurements
  Conclusion and Outlook
    Limitations for BPM in circular Machine

## Slide 4: Outline
Electro-Optical Bunch Arrival-Time Measurement
Pickups for Bunch Arrival-Time Monitors
  From button-pickups to 40-GHz cone-shaped pickups
Pickups on a Printed Circuit Board (PCB)
  Concept for a PCB-BAM
  Production of a first demonstrator
  Measurements
  Conclusion and Outlook
    Limitations for BPM in circular Machine

## Slide 5: Electro-Optical Bunch Arrival-Time Measurement
RF-part
Pickup-structure at EuXFEL (left) & ELBE (right)
Key Components:
Simplified EO-BAM layout with only one pickup and one EOM
40-GHz cone-shaped pickup [Angelovski 2012]
Exemplary pickup signal
A. Angelovski et al., Proc. IBIC’12. Paper: MOPA46
Transmittance of an idealized EOM vs signal voltage
Diode
Electronics
EOM
fs pulsed
optical
reference
Electro-optical unit
Data acquisition

## Slide 6: Electro-Optical Bunch Arrival-Time Measurement
EO-BAM – Detection scheme
Reference: Laser pulse samples signal @ zero crossing (ZC)
Ideal timing: Output amplitude same as the input amplitude
Early bunch: Lower output amplitude
Late bunch: Higher output amplitude
Laser intensity proportional to
All are linear for small deviations
RF Input
Pickup Voltage Signal
F. Löhl et al., Phys. Rev. Lett, 2010. DOI:10.1103/PhysRevLett.104.144801
Measurement Principle [based on Löhl, 2010]
Diode
fs pulsed optical reference
DC bias
EOM
Input Voltage
Intensity
Output
Notes: Sensititvity comes with reduced dynamic range

## Slide 7: Electro-Optical Bunch Arrival-Time Measurement
Comparison to a Button-BPM
 | Typical EO-BAM | Typical Button-BPM
Pickups | Cone-shaped (button-like) | Buttons
No. of PU | 2 - 4 (1 - 2 Channels) | 4 (vert. & horz. channel)
Termination |  | 
Cut-off |  | MHz to lower GHz
Signal Type | Derivative | Derivative or Proportional
 | Typical EO-BAM | Typical Button-BPM
Pickups | Cone-shaped (button-like) | Buttons
No. of PU | 2 - 4 (1 - 2 Channels) | 4 (vert. & horz. channel)
Termination |  | 
Cut-off |  | MHz to lower GHz
Signal Type | Derivative | Derivative or Proportional

## Slide 8: Electro-Optical Bunch Arrival-Time Measurement
Comparison to a Button-BPM
 | Typical EO-BAM | Typical Button-BPM
Pickups | Cone-shaped (button-like) | Buttons
No. of PU | 2 - 4 (1 - 2 Channels) | 4 (vert. & horz. channel)
Termination |  | 
Cut-off |  | MHz to lower GHz
Signal Type | Derivative | Derivative or Proportional
 | Typical EO-BAM | Typical Button-BPM
Pickups | Cone-shaped (button-like) | Buttons
No. of PU | 2 - 4 (1 - 2 Channels) | 4 (vert. & horz. channel)
Termination |  | 
Cut-off |  | MHz to lower GHz
Signal Type | Derivative | Derivative or Proportional
A. Angelovski et al., Proc. IBIC’12. Paper: MOPA46
P. Forck, Proc. CAS 2024, Event:1356988, Contr.:5713247
Angelovski, IBIC’12
hybrid
diff
sum
Extract of an exemplary broadband signal processing scheme for a button-BPM [adapted from P. Forck, 2024].
Green parts are also common for the EO-BAM.
Two different problems, similar solutions!

## Slide 9: Outline
Electro-Optical Bunch Arrival-Time Measurement
Pickups for Bunch Arrival-Time Monitors
  From button-pickups to 40-GHz cone-shaped pickups
Pickups on a Printed Circuit Board (PCB)
  Concept for a PCB-BAM
  Production of a first demonstrator
  Measurements
  Conclusion and Outlook
    Limitations for BPM in circular Machine

## Slide 10: Pickups for Bunch Arrival-Time Monitors
Button-type pickups:
Bandwidth: Pickup ; EOM
Bandwidth Showed viability of EO-BAM measurement:
Strong ringing
2009
F. Löhl & K. Hacker
until
3D model of the button-type beam pickup designed by Hacker and used for the EO-BAM by Löhl [Löhl 2009]
F. Löhl , Thesis, 2009. DOI:10.3204/DESY-THESIS-2009-031
Technical drawing and 3D model of the button-type beam pickup de-signed by K. Hacker and used for the EO-BAM by F. Löhl [Löhl 2009].
Measured signal of two combined pickups with an 8-GHz oscilloscope [Löhl 2009]

## Slide 11: Pickups for Bunch Arrival-Time Monitors
Matching:
Goal: Reduction of ringing!
Using TEM wave model:
  only the TEM wave can propagate (at all frequencies)
  no other waves with axial components (TMmn and TEmn waves) should be propagating
  TE11 lowest cut-off → Defines of a coaxial:
Line Impedance:
2009
F. Löhl & K. Hacker
until
3D model of the button-type beam pickup designed by Hacker and used for the EO-BAM by Löhl [Löhl 2009]
F. Löhl , Thesis, 2009. DOI:10.3204/DESY-THESIS-2009-031
Z0: Free-Space wave impedance
Field distribution of TEM (left) an unwanted TE11 mode (right) in a coaxial cable
Cross section of a coaxial cable
Inner conductor
outer conductor
Dielectric, e.g., air
Notes: Epsilon_R = 10  13.95
Epsilon_R = 3  4.23

## Slide 12: Pickups for Bunch Arrival-Time Monitors
Matching:
Design goal: ;
Ratio D/d: ▬▬►
: and
The radii must be smaller than defined above, to allow for TEM only propagation up to
2009
F. Löhl & K. Hacker
until
3D model of the button-type beam pickup designed by Hacker and used for the EO-BAM by Löhl [Löhl 2009]
F. Löhl , Thesis, 2009. DOI:10.3204/DESY-THESIS-2009-031
(in air!)
Z0: Free-Space wave impedance
Field distribution of TEM (left) an unwanted TE11 mode (right) in a coaxial cable
Cross section of a coaxial cable
Inner conductor
outer conductor
Dielectric, e.g., air

## Slide 13: Pickups for Bunch Arrival-Time Monitors
1st generation of cone-shaped pickups:
Increased rf bandwidth of all components ()
Avoids structural resonances by gaps, discontinuities & sharp edges
Avoids reflections by maintaining geometry at every cross-section
2009
F. Löhl & K. Hacker
until
3D model of the button-type beam pickup designed by Hacker and used for the EO-BAM by Löhl [Löhl 2009]
F. Löhl , Thesis, 2009. DOI:10.3204/DESY-THESIS-2009-031
A. Angelovski et al., Proc. IPAC’11, Paper: TUPC076
2011
A. Angelovski
Cross-section of the cone-shaped pickup [Angelovski 2011]
Cross-section of the cone-shaped pickup [Angelovski 2011]
Signal output simulated with CST PS for a 1-mm 20-pC bunch [Angelovski 12]
A. Angelovski et al., Phys. Rev. ST Accel. Beams, 2012. DOI:10.1103/PhysRevSTAB.15.112803
Notes: # Reduced capacitance (smaller time constant)
# Tapered cut-out with constant ratio b / a = 2.3 for 50 Ω
# (Nearly) Resonance-free up to 40 GHz

## Slide 14: Pickups for Bunch Arrival-Time Monitors
1st generation of cone-shaped pickups:
Vacuum sealed prototype successfully tested
But: Better agreement with simulation for non-hermetic demonstrator
2009
F. Löhl & K. Hacker
until
3D model of the button-type beam pickup designed by Hacker and used for the EO-BAM by Löhl [Löhl 2009]
F. Löhl , Thesis, 2009. DOI:10.3204/DESY-THESIS-2009-031
A. Angelovski et al., Proc. IPAC’11, Paper: TUPC076
2011
A. Angelovski
Cross-section of the cone-shaped pickup [Angelovski 2011]
A. Angelovski et al, in Proc IBIC2012, Tsukuba, Japan, 2012
Pickup top view
Pickup side view
Hermetic demonstrator of the 1st generation cone-shaped pickup [Angelovski 12]
Vacuum flange
(Caused by step
in the beamline)
Notes: # Reduced capacitance (smaller time constant)
# Tapered cut-out with constant ratio b / a = 2.3 for 50 Ω
# (Nearly) Resonance-free up to 40 GHz

## Slide 15: Pickups for Bunch Arrival-Time Monitors
2nd generation of cone-shaped pickups:
1st generation did not reach sufficient voltages
Cone diameter: vs. trade-off
2nd gen. optimized for maximum voltage maintaining a slope of at least
2009
F. Löhl & K. Hacker
until
3D model of the button-type beam pickup designed by Hacker and used for the EO-BAM by Löhl [Löhl 2009]
F. Löhl , Thesis, 2009. DOI:10.3204/DESY-THESIS-2009-031
A. Angelovski et al., Proc. IPAC’11, Paper: TUPC076
A. Angelovski et al., Proc. IBIC’13, Paper: WEPC40
2011
A. Angelovski
Cross-section of the cone-shaped pickup [Angelovski 2011]
2013
Cross section of the cone-shaped pickup (left) and the modified cone-shaped pickup for high (right) [Angelovski 2013]
A. Angelovski
Cross section of cone-shaped pickup (left) & modified cone-shaped pickup for high (right) [Angelovski 2013]
and SR at ZC vs. cone diameter. The optimum area is shaded [Angelovski 13].
Cone Diameter / mm

## Slide 16: Pickups for Bunch Arrival-Time Monitors
2nd generation of cone-shaped pickups:
1st generation did not reach sufficient voltages
Cone diameter: vs. trade-off
2nd gen. optimized for maximum voltage maintaining a slope of at least
increased by a factor of 2
But: Increased cross talk
2009
F. Löhl & K. Hacker
until
3D model of the button-type beam pickup designed by Hacker and used for the EO-BAM by Löhl [Löhl 2009]
F. Löhl , Thesis, 2009. DOI:10.3204/DESY-THESIS-2009-031
A. Angelovski et al., Proc. IPAC’11, Paper: TUPC076
A. Angelovski et al., Proc. IBIC’13, Paper: WEPC40
2011
A. Angelovski
Cross-section of the cone-shaped pickup [Angelovski 2011]
2013
Cross section of the cone-shaped pickup (left) and the modified cone-shaped pickup for high (right) [Angelovski 2013]
A. Angelovski
Comparison of simulated signals & crosstalk (S21) for 1st and 2nd gen. pickups [Angelovski 13].

## Slide 17: Pickups for Bunch Arrival-Time Monitors
Broadband cone-shaped pickups:
Commercialized by DESY
Installed at:
2009
F. Löhl & K. Hacker
until
3D model of the button-type beam pickup designed by Hacker and used for the EO-BAM by Löhl [Löhl 2009]
F. Löhl , Thesis, 2009. DOI:10.3204/DESY-THESIS-2009-031
A. Angelovski et al., Proc. IPAC’11, Paper: TUPC076
A. Angelovski et al., Proc. IBIC’13, Paper: WEPC40
2011
A. Angelovski
Cross-section of the cone-shaped pickup [Angelovski 2011]
2013
Cross section of the cone-shaped pickup (left) and the modified cone-shaped pickup for high (right) [Angelovski 2013]
A. Angelovski

## Slide 18: Outline
Electro-Optical Bunch Arrival-Time Measurement
Pickups for Bunch Arrival-Time Monitors
  From button-pickups to 40-GHz cone-shaped pickups
Pickups on a Printed Circuit Board (PCB)
  Concept for a PCB-BAM
  Production of a first demonstrator
  Measurements
  Conclusion and Outlook
    Limitations for BPM in circular Machine

## Slide 19: Pickups on a Printed Circuit Board
Concept for a PCB-BAM:
First promising concept deviated from the PCB idea
Reached 80 % of the project goal
Ringing and slew rate strongly depending on protrusion length
2021
B.E.J. Scheible
Open-coaxial pickups with elongated pins [Scheible 2021]
B.E.J. Scheible et al., Proc. IBIC’21. 10.18429/JACoW-IBIC2021-WEPP19
Open-coaxial pickups with elongated pins and a tip-to-tip distance of 10 mm [Scheible 2021]
@ ,
Pickup signal simulated by CST PS for a very short and a relatively long rod length [Scheible 2021]

## Slide 20: Pickups on a Printed Circuit Board
Concept for a PCB-BAM:
Use printed circuit boards with planar transmission line
Easiest type: Microstrip line (quasi-TEM if sufficiently thin)
2021
B.E.J. Scheible
Open-coaxial pickups with elongated pins [Scheible 2021]
B.E.J. Scheible et al., Proc. IBIC’21. 10.18429/JACoW-IBIC2021-WEPP19
Electric & magnetic fields on a microstrip [Pozar 2012]
D. Pozar, “Microwave Engineering”, 4th Ed., 2012, 978-0-470-63155-3 F. De Flaviis, in The Electrical Engineering Handbook, 2005, downloaded from sciencedirect.com/topics/computer-science/microstrip at 09.12.2024
Strip Conductor
Conductor (Ground plane)
Dielectric substrate
Geometry of a microstrip line [De Flaviis 2005]

## Slide 21: Pickups on a Printed Circuit Board
Concept for a PCB-BAM:
Followed by a combination of former concepts (PCB & open coax) with integrated signal combination on the PCB
First to surpass project goal
But: Ringing and wake fields
2021
B.E.J. Scheible
Open-coaxial pickups with elongated pins [Scheible 2021]
B.E.J. Scheible et al., Proc. IBIC’21. 10.18429/JACoW-IBIC2021-WEPP19
2021
B.E.J. Scheible
Open-“coax“ pickups on a PCB with Ø 10 mm opening [Scheible 2021]
Open-“coax“ pickups on a PCB with Ø 10 mm opening [Scheible 2021]
@ ,
Pickup signal simulated by CST PS for a 4-rod on PCB pickup structure with combination network [Scheible 2021]

## Slide 22: Pickups on a Printed Circuit Board
Concept for a PCB-BAM:
Final concept: Planar elongated pickup on a reduced PCB with integrated combination network
Reduced Wakefield effects
Easier to manufacture
2021
B.E.J. Scheible
Open-coaxial pickups with elongated pins [Scheible 2021]
B.E.J. Scheible et al., Proc. IBIC’21. 10.18429/JACoW-IBIC2021-WEPP19
B.E.J. Scheible et al., Proc. IPAC’23. 10.18429/JACoW-IPAC2023-THPL190
2021
B.E.J. Scheible
Open-“coax“ pickups on a PCB with Ø 10 mm opening [Scheible 2021]
2023
Planar pickups on a printed circuit board [Scheible 2023]
B.E.J. Scheible
Planar pickups on a printed circuit board [Scheible 2023]
@ ,
Pickup signal simulated by CST PS for planar pickups on PCB with combination network [Scheible 2023]

## Slide 23: Outline
Electro-Optical Bunch Arrival-Time Measurement
Pickups for Bunch Arrival-Time Monitors
  From button-pickups to 40-GHz cone-shaped pickups
Pickups on a Printed Circuit Board (PCB)
  Concept for a PCB-BAM
  Production of a first demonstrator
  Measurements
  Conclusion and Outlook
    Limitations for BPM in circular Machine

## Slide 24: Pickups on a Printed Circuit Board
Production of a first demonstrator
Rapid prototyping at THM
Design of vacuum-sealed demonstrator
2021
B.E.J. Scheible
Open-coaxial pickups with elongated pins [Scheible 2021]
B.E.J. Scheible et al., Proc. IBIC’21. 10.18429/JACoW-IBIC2021-WEPP19
B.E.J. Scheible et al., Proc. IPAC’23. 10.18429/JACoW-IPAC2023-THPL190
2021
B.E.J. Scheible
Open-“coax“ pickups on a PCB with Ø 10 mm opening [Scheible 2021]
2023
Planar pickups on a printed circuit board [Scheible 2023]
B.E.J. Scheible
dowel pin
dowel pin
stopper
67-GHz connector
PCB
10 cm
sheet pickups
DN25 CF
DN40 CF
3.6 cm
feedthrough module
Manufacturing model (in Inventor®) of the first vacuum-sealed prototype for ELBE and steel body built by NTG

## Slide 25: Pickups on a Printed Circuit Board
Production of a first demonstrator
Rapid prototyping at THM
2021
B.E.J. Scheible
Open-coaxial pickups with elongated pins [Scheible 2021]
B.E.J. Scheible et al., Proc. IBIC’21. 10.18429/JACoW-IBIC2021-WEPP19
B.E.J. Scheible et al., Proc. IPAC’23. 10.18429/JACoW-IPAC2023-THPL190
2021
B.E.J. Scheible
Open-“coax“ pickups on a PCB with Ø 10 mm opening [Scheible 2021]
2023
Planar pickups on a printed circuit board [Scheible 2023]
B.E.J. Scheible
Sample PCB with 2 € coin for scale
RF PCB milling machine at THM
Broken spare PCB made with TMM10i

## Slide 26: Pickups on a Printed Circuit Board
Production of a first demonstrator – Vacuum feedthroughs
Many connectors commercially available (even above 100 𝐺𝐻𝑧)
Typically, insufficient hermeticity and not well-suited for accelerators
Picture of a Rosenberger right angle jack for PCBs and drawing from the technical datasheet
rosenberger.com, Rosenberger No. 01K80A-40ML5, downloaded Nov. 28th, 2024
anritsu.com, Anritsu V1-102F Female Connector with V100 glass bead , downloaded Nov. 28th, 2024
Not suited as vacuum feedthrough !
Hermetic* and can be used as vacuum feedthrough !
Picture of an Anritsu female sparkplug launcher connector and drawing from the instruction sheets
*HERMETIC TO 1X10-8 STD CC HE/SEC AT 1 ATM DIFFERENTIAL.

## Slide 27: Pickups on a Printed Circuit Board
Production of a first demonstrator – Vacuum feedthroughs
Many connectors commercially available (even above 100 𝐺𝐻𝑧)
Typically, insufficient hermeticity and not well-suited for accelerators
Some connectors can be used as a vacuum feedthrough!
Installation still is tricky!
rosenberger.com, Rosenberger No. 01K80A-40ML5, downloaded Nov. 28th, 2024
anritsu.com, Anritsu V1-102F Female Connector with V100 glass bead , downloaded Nov. 28th, 2024
A. Angelovski et al., Proc. IPAC’11, Paper: TUPC076
Cross-section of the cone-shaped pickup [Angelovski 2011] and picture of an Anritsu female sparkplug launcher connector and drawing from the instruction sheets
Vacuum feedthrough module

## Slide 28: Pickups on a Printed Circuit Board
Production of a first demonstrator – Vacuum feedthroughs (Non hermetic demonstrator)
No Undo-Button!

## Slide 29: Pickups on a Printed Circuit Board
Production of a first demonstrator – Mounting of the PCB
Bead with sliding contact & sliding contact on a single wire
Sliding contact on the PCB
One day of failed approaches…
Positioning with 3D stage (successful)
Broken spare of the PCB
Backside of the PCB resting on the ledge
PCB through the DN25 CF (left) respectively DN40 CF flange (center) and front of the PCB before mounting the feedthrough and locking mechanism (opposite side).
Rogers
TMM10i

## Slide 30: Outline
Electro-Optical Bunch Arrival-Time Measurement
Pickups for Bunch Arrival-Time Monitors
  From button-pickups to 40-GHz cone-shaped pickups
Pickups on a Printed Circuit Board (PCB)
  Concept for a PCB-BAM
  Production of a first demonstrator
  Measurements
  Conclusion and Outlook
    Limitations for BPM in circular Machine

## Slide 31: Pickups on a Printed Circuit Board
Measurement with the first demonstrator
Proof-of-concept measurements at the ELBE accelerator
Signal similar to simulation
First pulse ever measured with the EO-PCB BAM (Acq002-CH1-Pulse0) with an expected charge of at HZDR in 2023

## Slide 32: Pickups on a Printed Circuit Board
Measurement with the first demonstrator
Proof-of-concept measurements at the ELBE accelerator
Signal similar to simulation
Slew-rate below project goal, mostly due to choice of substrate
Signal slew rate (SR) vs charge for the first demonstrator compared to different simulations and the current cone-shaped pickups.
A. Angelovski et al., Phys. Rev. ST Accel. Beams, 2015. DOI:10.1103/PhysRevSTAB.18.012801
Decrease due to use of the TMM10i laminate
Sample of 0.2 mm fused-silica waver after laser-cutting at THM

## Slide 33: Outline
Electro-Optical Bunch Arrival-Time Measurement
Pickups for Bunch Arrival-Time Monitors
  From button-pickups to 40-GHz cone-shaped pickups
Pickups on a Printed Circuit Board (PCB)
  Concept for a PCB-BAM
  Production of a first demonstrator
  Measurements
  Conclusion and Outlook
    Limitations for BPM in circular Machine

## Slide 34: Pickups on a Printed Circuit Board
Conclusion
Planar pickups on a printed circuit board give a strong signal
Well suited for short bunches and “low” repetition rates
Drawbacks:
  High wake effect due to discontinuity in beamline
  Strong ringing after main pulse
PCB is versatile:
  Subtraction of two signals should be possible, but ultra-wideband will be a challenge!
Broken spare of the PCB
Angelovski, IBIC’12
hybrid
Extract of an exemplary broadband signal processing scheme for a button-BPM [adapted from P. Forck, 2024]
P. Forck, Proc. CAS 2024, Event:1356988, Contr.:5713247
Additional differential path?

## Slide 35
Further Questions?
Bernhard E. J. Scheible
bernhard.scheible@iem.thm.de
Technische Hochschule Mittelhessen
Information Technology | Electrical Engineering | Mechatronics
This work is supported by the German Federal Ministry of Education and Research (BMBF) under contract no. 05K22RO2.

## Slide 36: APPENDIX: Achieved BAM resolution
European XFEL reached at nominal [Czwalinna, 2021]
SwissFEL reached at nominal due to smaller beampipe [Arsov, 2024]
Resolution depends on bunch charge and beampipe radius (etc.)
M. K. Czwalinna et al., Proc. IPAC’21. DOI:10.18429/JACoW-IPAC2021-THXB02
V. Arsov et al., Proc. IPAC’24 (pre press). DOI:10.18429/JACoW-IPAC2024-WEPG68
Single-bunch correlation for 1 min of data [Czwalinna 2021]
Charge dependence of the BAM resolution [Arsov 2024]
Notes: They [BAM3 & BAM4] are separated by ca. 1.5 km straight section of the linac without any significant longitudinally dispersive element in between. [Czw.21]

## Slide 37: APPENDIX: Expected Resolution
Estimated for a system with…
rod-shaped pickups on PCB with integrated combination (simulation)
silicon-organic hybrid MZM (measured demonstrator)
20 cm coaxial cable (model)
Different sources of jitter (typical, literature, …)
Theoretical jitter-charge product:
Dominated by relative laser intensity noise -> Balanced design
U (V)
t (ps)
Notes: RIN Nois Laser, q Quantum shot noise, pd Photodiode, ADC ADC Digitizer

## Slide 38: APPENDIX: Charge Estimation and Error
B.E.J. Scheible et al., Proc. IBIC’23. DOI:10.18429/JACoW-IBIC2023-TUP012
B.E.J. Scheible et al., Proc. IPAC’24. DOI:10.18429/JACoW-IPAC2024-WEPG82
Charge not known exactly (ICT below resolution limit), but
Approximation of bunch charge:
Error estimation:
Example:
Error dominated by the resolution of the cathode current! Error of bunch spacing is negligible.
Example values:
,


## Slide 39: APPENDIX: ELBE Parameter and Layout
ELBE Facility Layout [Source: hzdr.de]
Position of the PCB-BAM Demonstrator
Parameter
Thermionic gun
CW mode
Repetition rate:
Current:
Charge:
Bunch length:
PCB-BAM in Beamline

## Slide 40: APPENDIX: Pre-Fields
Time travelling along transmission line

## Slide 41: APPENDIX: Amplitude Jitter

## Slide 42: APPENDIX: Optical Link Stabilization
Laser reference and stabilized distribution
MLO phase-locked to MO
Reference pulse split in free-space
Distribution in actively stabilized fiber links
  Links up to >3500 m
  Optical-cross correlator (OXC)
  Piezo fiber stretcher & optical delay line (ODL)
Link stabilization unit on optical the table in the European XFEL
Optical table with MLO and OXC in the European XFEL
MLO
fs-pulsed optical reference
OXC
detection
piezo
mech. ODL
correction
up to 3.5 km
long-haul-fiber link
sub-fs jitter
client system
partial reflection
link end
Layout of one active-stabilized fiber link at the European XFEL [adapted from S. Schulz, ARD-ST3 Workshop 2023]
Notes: Ref: Sydlo 2015
Optical tables made of Superinvar (low thermal expansion), thermal stability of < 0.1 K
Split for up to 24 links
Link length: 35 m to >3500 m
Delay-Line: 4 ns, slow
Fiber Stretcher: ~3.5 ps, fast correction
First P-Waves, later S-wave -> Very Sensitve system!

## Slide 43: APPENDIX: Analytical Signal
Long Bunch Approximation:
is constant in any interval of length
e-
Independent of the pickup shape!

## Slide 44: APPENDIX: Analytical Signal
Short Bunch Approximation:
For any , which satisfies the requirements of a Dirac sequence:
e-
Independent of the bunch shape!