## Slide 1: Beamline Control and Data Acquisition System at SSRF
Zhang Zhaohong, Zheng Lifang
Beamline Control Group
Shanghai Synchrotron Radiation Facility
about 25 pages

## Slide 2: Outlines
Overview
Infrastructure
Control System
Data Acquisition System
Current Missions
Conclusion
Average 4 slides per section

## Slide 3: Overview
1.1 Organization
1.2 Responsibilities
1.3 System Overview

## Slide 4: 1.1 Organization
Beamline Engineering Department
Experiment Departments
IT Department
Control Group
Electronic & Detector
Group
13 Beamlines
1 group leader
7 engineers
1 group leader
5 engineers
1 technician
Each scientific village has 1 to 3 control engineers
~ 4 related people

## Slide 5: 1.2 Responsibilities
Control Group
E & D Group
IT Dep.
Experiment Dep.
Beamline Control
Personnel Safety
Equipment Protection
Endstation Control
Data Acquisition
Endstation Control (participate )
Network
Computing
& storage
Beam Monitoring
infrastructure
Control group cooperates with beamline staff to complete the whole control and data acquisition system

## Slide 6: 1.3 System overview
EPICS
EPICS: Experimental physical industrial control system
EPICS CA protocol
Based on TCP/IP
Platform
Control system
Communication Protocol
Computer
interface
PLC system
Personnel safety
PLC system
Equipment protection
DAS
Labview
CA interface
DAS
VC++
DAS
Blu-Ice
DAS
SPEC
DAS
EPICS
DAS: Data acquisition system

## Slide 7: 2. Infrastructure
2.1 Network
2.2 Personnel Safety Interlock System
2.3 Equipment Protection System

## Slide 8: 2.1 Network Structure
SSRF
Central Switch
Beamline
Central Switch
Control Room
Switch
Other Beamline
Switches
MX & Image beamline Switches
Each beamline has an independent subnet.
Computing Center
Cluster Switch
In the future
2

## Slide 9: 2.2 Personnel Safety Interlock System
Digital Values
Photon Shutter
Safety Shutter
Search Buttons
Door Switches
Emergency buttons
Interlock key
PLC Modules
Get digital values
Calculate logics
Set up interlocks
Give alarming signal
EPICS software
Have the capacity to access each tag in PLC
PLC Bus
Ethernet
Search Buttons
Control Panel &
Display Screen
A-B
Control-Logix 5561
DI
DO

## Slide 10: 2.3 Equipment Protection System
PLC modules
PLC modules
Water Flow state
Temperature
Vacuum
PLC system
Ethernet
Serial Server
EPICS Software
etc.
PLC touch panels
OPI
Vacuum Gauge

## Slide 11: 3 Control System
3.1 EPICS Overview
3.2 Motion Control
3.3 Message Based Devices’ Control
3.4 Detector Control
3.5 Scan Mechanism
3.6 OPI Tools
3.7 Other Tools
Altogether 10 slides

## Slide 12: 3.1 EPICS Overview
We benefit a lot from EPICS, developed so perfectly.
OPI Layer
LAN: CA Protocol
VME-IOC
PC-IOC
PC-IOC
Motion Control
status monitor
Detector
Control
Data
Acquisition
EPICS is competent for any control task in beamlines.

## Slide 13: 3 Control System
3.1 EPICS Overview
3.2 Motion Control
3.3 Message Based Devices’ Control
3.4 Detector Control
3.5 Scan Mechanism
3.6 OPI Tools
3.7 Other Tools
Altogether 10 slides

## Slide 14: 3.2.1 Mathematical Model
Control Objects
Slit
Mirror Chamber
Monochromator
Sample Stage
Filter
BPM
etc.
Execute Components
step motors (over 95%)
Control Variables
Size,offset
Height,pitch,roll
Energy, etc.
…
…
…
…
=
multi-input & multi-output
linear system
Beam line is all about motors
Motion control is everywhere.
Position control !

## Slide 15: 3.2.2 EPICS facilities
=
Output
Algorithm
Input
In EPICS’s world, everything is PV. PVs = fields of EPICS record .
database
records
PVS
EPICS IOC
Low level drivers
OPI
motor record
~100 fields
softMotor record
~100 fields
Transform record
A~P, float fields
CLC field(C language calculation)
record Links
INP OUT types
PP NPP CP modes
We formed a Standardized method to implement motion systems
using the elements in EPICS.

## Slide 16: 3.2.3 example
Controlled variables
motors
softMotor records
motor records
transform records
Var. | motors and formulas
y | y=(2*y1+y2+y3)/4
pitch | Pitch=atg (((y2+y3)-2y1)/2)
roll | Roll=atg ((y3-y2)/W)
x | x=(x1+x2)/2
yaw | yaw=atg ((x2-x1)/L)
Implement the formulas in IOC

## Slide 17: 3.2.4 hardware for step motor control
8 axes
Open/closed loop
EPICS driver support
VME single board computer MVME 5500
CPU: 1GHz
On-board memory: 512M
Ethernet interface: gigabit
Bus speed: 66MHz
OS: VxWorks 5.5
Custom-build standard power
rack & modular motor drivers
Power supply: 24 VDC/48 VDC
2-phase, 5-phase
0.2~6A driver current
0~100% holding current
Phytron’s core motor power stage
Normalized connector and cabling discipline
MAXv-8000
motion controller
Standardized & normalized hardware
High efficiency & Good result

## Slide 18: 3 Control System
3.1 EPICS Overview
3.2 Motion Control
3.3 Message Based Devices’ Control
3.4 Detector Control
3.5 Scan Mechanism
3.6 OPI Tools
3.7 Other Tools
Altogether 10 slides

## Slide 19: 3.3 Message Based Devices
Epics communicates with devices through serial port or
ethernet port.
RS-232
RS-485
Ethernet
Serial
Server
EPICS
Soft IOC
LAN
streamDevice
devGPIB
Asyn
Configure a protocol file
Configure a C struct
Write some C codes
Three independent methods, the more complicated, the more freedom
Simple
Complicated
PI-E621
PI-C863
Newport-XPS
Pro-dex MAXnet

## Slide 20: 3 Control System
3.1 EPICS Overview
3.2 Motion Control
3.3 Message Based Devices’ Control
3.4 Detector Control
3.5 Scan Mechanism
3.6 OPI Tools
3.7 Other Tools
Altogether 10 slides

## Slide 21: 3.4 Detector Contol
Ionization Chambers
Solid Detectors
CCDS
picoammeter + EPICS
VME ADC card + EPICS
V/F + VME scalar + EPICS
NI ADC + Labview
XIA electronic readout system
& EPICS
EPICS & image software
Thank EPICS community!
There are so much hardware
used in the worldwide
synchrotron facilities with EPICS.
It’s a shortcut to choose
among these hardware.
We share the benefit of
EPICS/synApps package, which has many support modules for specific devices.

## Slide 22: 3 Control System
3.1 EPICS Overview
3.2 Motion Control
3.3 Message Based Devices’ Control
3.4 Detector Control
3.5 Scan Mechanism
3.6 OPI Tools
3.7 Other Tools
Altogether 10 slides

## Slide 23: Scan Mechanism
Start scan
Set scan.EXCS = 1
Move to the next desired point
Wait DLY seconds for positioners to settle down
Trigger the detector
Wait DDLY seconds for detectors to settle dpwn
Stop scan
Continuous loops until the last point
Some predefined functions can be invoked during these periods.
Scan mechanism is the most important
step by step scan
SSCAN soft modules
  Scan record
  SaveData function
  Plotting function
Every numeric PV is scannable
  motor.VAL
Every numeric PV can act as a detector.
Data of positioners and detectors are collected in scan record’s array fields.
1-D、2-D and multi-D scan can be achieved flexibly.
Example:
Do averaging
Post data

## Slide 24: 3 Control System
3.1 EPICS Overview
3.2 Motion Control
3.3 Message Based Devices’ Control
3.4 Detector Control
3.5 Scan Mechanism
3.6 OPI Tools
3.7 Other Tools
Altogether 10 slides

## Slide 25: 3.6 OPI Tools
EDM @ Linux
CSS@Linux&windows
medm @ Linux&windows
EDM: Extemsible Display Manager
MEDM: Motif Editor and Display Manager
CSS: Control System Studio
Probe
plot2D
stripTool
PV finder
dataBrowser
PV tree
Static graph
widgets
Monitoring
widgets
control
widgets
probe

## Slide 27: 3. Control System
3.1 EPICS Overview
3.2 Motion Control
3.3 Message Based Devices’ Control
3.4 Detector Control
3.5 Scan Mechanism
3.6 OPI Tools
3.7 Other Tools
Altogether 10 slides

## Slide 28: 3.7 Other Tools
EPICS Sequencer
  one component of EPICS IOC
  access/compute /analyse
  C or SNL language
  example: FFT; event trigger generator
EPICS Archiver
  one EPICS tool
  archive PVS
  plot historic data
Python\Jython\Javascript \IDL
  with CA interface to EPICS
  flexibly do high level applications
Pviewer
  motif tool to manage the .MDA data saved by the scan record
imageJ / fid2d / igor / origion
  tools to show and analyse image data or 2D ASC-II data

## Slide 29: 4. Data Acquisition System
4.1 Overview
4.2 Example: Blu-Ice

## Slide 30: 4.1 DAQ Overview
Beamline Name | Beamline Control | Endstation Control | Data Acquisition
XAFS | EPICS | EPICS | LabView
STMX | EPICS | 3rd party software | VC++
XIM | EPICS | LabView | VC++
XRD | EPICS | 3rd party software | SPEC
SAXS | EPICS | EPICS | EPICS
XMF | EPICS | EPICS | EPICS
5 MX Beamlines | EPICS | EPICS 3rd party software | BlueIce
Dreamline | EPICS | 3rd party software | 3rd party software
Infrared Beamline | EPICS | 3rd party software | 3rd party software

## Slide 31: 4.2 Example: Blu-Ice
EPICS
Control System
DHS
TangoDeviceServer
Call API
EPICS gateway
DHS
CameraMan
CamServer
DHS
DCSS
DCSS: Distributed Control System Server
DHS: Distributed Hardware Server
GUI
LAN
GUI: Graphic User Interface
servers
TK

## Slide 32: 5. Current Missions
5.1 Time Resolved Experiments
5.2 Sub-second QXAFS
5.3 Feedback System

## Slide 33: 5.1 Time-resolved Experiments
Laser Pump-probe system
Pump: Laser pulse
Probe: X-ray pulse
To study the intermediate state of the reversible reaction
Resolution: ~100 ps
Time structure of a special electron bunch pattern
TR-XEOL
Pump: single X-ray pulse
Probe: Electronic devices
To get the luminescence decay curve
Resolution: 2ns
XEOL: X-ray Excited Optical Luminescence

## Slide 34: 5.2 Sub-second QXAFS
Key techniques:
synchronous triggers
timing system

## Slide 35: 5.3 Feedback System
To develop a universal high performance digital feedback controller based on FPGA
PSD
BBPM
CCD
photodiodes
CCD or other devices
Laser interferometer
Many requirements in our coming beamlins
Beam
Position
Feedback
Beam
Intensity
Feedback
Nano-meter focusing
Feedback
Monochromator
Mirrors
Nano-meter sample stage

## Slide 36: 6. Conclusion

## Slide 37: 6. Conclusion
Standard control system has been developed
  Standard hardwares
  standard software based on EPICS
  EPICS does a lot for us, we should do more for EPICS next.
Data acquisiton systems are various
  Wish to develop a universal one
  With generic application framework, common feel&like GUIs, configurable and reusable components, capacity to trace behaviors, and “learn once, use everywhere” features.

## Slide 38: 6. Conclusion
Great challenges are waiting for us
  picosecond time-resolved experiments.
  nano-meter feedback controls
  sub-micrometer beam detect technologies
  information management system based on database
We hope to collaborate with experts
throughout the world.

## Slide 39
Thank you for your attention