## Slide 1: Streak camera optimisation at Diamond
L. M. Bobb and G. Rehm

## Slide 2: Overview
PROBLEM | SOLUTION
Low alpha beam mode: Single bunch investigations Low beam current (~µA) Low flux at streak camera High resolution required to measure short bunch lengths (1ps) | Reflective input optics  Boost flux Pinhole  Obtain small, round PSF Calibration of dual sweep time bases  High resolution measurements
Streak Camera (SC):
Optronis Optoscope SE, FSSU1-H2, STSU2
PSF = Point Spread Function
= Spot size with dual sweeps disabled

## Slide 3: Input Optics Progression (1)
SC
Astigmatism
 Poor slow axis resolution
Breadboard optics:
System of focussing mirrors
= High Flux
Concave parabolic
Convex spherical
Concave spherical
Flat fold mirror
1
2
3
4

## Slide 4: Input Optics Progression (2)
SC
Pinhole
Lens
Fold mirror
Breadboard optics + filter
= Limited Flux
Imaging pinhole
= Small round PSF
Bandpass filter

## Slide 5: Input Optics Progression (3)
SC
Breadboard optics + 2x 90° OAP
= High Flux
Imaging pinhole
= Small round PSF
Pinhole
OAP
Fold mirror

## Slide 6: Electrical Delay Calibration Setup
For fast sweep calibration:
Difference in time between the 2 cables = 15.7 ps
For slow sweep calibration
Mini-circuits switches

## Slide 7: Calibration Results
Resolution setting [ps/mm] | Measured resolution [ps/pixel] | Which corresponds to… [ps/mm]
15 | 0.1816 (Feb 2015) 0.1998 (May 2015) | 12.97 14.27
25 | 0.3195 | 22.82
50 | 0.6466 | 46.19
Calibration
Difference between centroid positions in number of pixels
= 15.7 ps

## Slide 8: Performance
3.4 MV
alpha = +4.6×10-6
fs = 460 Hz
3.4 MV, alpha = -9.9×10-6, fs = 675 Hz, Ibunch = 95.9 μA
RECONSTRUCTION OF ELECTRON BUNCH MOTION DURING CSR BURSTS USING SYNCHRONISED DIAGNOSTICS
I.P.S. Martin et al.
Example of CSR bursting

## Slide 9: Thank you for your attention!

## Slide 10: Input Optics (1)

## Slide 11: Input Optics (2)

## Slide 12: Mini-Circuit Switches