## Slide 1: Status of Beam Loss Detector Tests and Developments at PETRA III & ESRF
Gero Kube
DESY (Hamburg)
Introduction: PETRA III Extension Project
Beam Loss Monitor Tests @ ESRF

## Slide 2: DESY Accelerator Complex
Gero Kube, DESY / MDI
DEELS 2015 Workshop @ ALBA, 15.06.2015
PETRA III
DESY II
LINAC II
PIA
FLASH I
FLASH II
E-XFEL
P3X-N
P3X-E

## Slide 3: PETRA III Extension Project
Gero Kube, DESY / MDI
circumference: 2304 m
energy: 6 GeV
emittance: 1 nm.rad
current: 100 mA
# bunches: 40 / 960
straight sections: 9
user beamlines: 14
DEELS 2015 Workshop @ ALBA, 15.06.2015
PETRA III parameters
Storage Ring PETRA III
→ 24
→ 1.2 nm.rad

## Slide 4: P3X Commissioning
Gero Kube, DESY / MDI
DEELS 2015 Workshop @ ALBA, 15.06.2015
project schedule
3.2.2015: first beam back in PETRA
13.2.2015: first stored beam
external users: end of April 2015
end of March 2015: back in user operation

## Slide 5: P3X: Diagnostics Tasks
Gero Kube, DESY / MDI
DEELS 2015 Workshop @ ALBA, 15.06.2015
expectation…
Ubaldo: „I'm sure you guys at Desy have interesting studies to show at the DEELS15!”
…and reality
how to operate sensitive electronics (Liberas) if only
the shell of the buiding is complete…
how to seal the roof of your diagnostics cabinet
if you have normal rain in Hamburg…
how to dry and clean the components
if it was not successful…
and so on…

## Slide 6: PETRA III: Touschek Factory…
Gero Kube, DESY / MDI
operation in time resolved (40 bunch) mode
Touschek limited lifetime
→ τ ≈ 1.4 h
observation of radiation damage
→ rust at undulator magnets
→ undulator performance degradation
→ damaged cables
top-up operation
DEELS 2015 Workshop @ ALBA, 15.06.2015
magnet structure of two undulators had to be replaced during P3X shutdown
PETRA III: need for new BLM system
talk Dirk Nölle (DESY): E-XFEL standard beam diagnostics
→ XFEL-BLM system based on PMT and radiator
DEELS 2014 @ ESRF
XFEL-BLMs for PETRA III ???

## Slide 7: BLM Tests
Gero Kube, DESY / MDI
P3X shutdown: no tests possible @ DESY
→ proposal: ESRF (Kees Scheidt)
DEELS 2015 Workshop @ ALBA, 15.06.2015
BLM installation @ ESRF in cell 4 behind scraper
BLM setup
Hamamatsu R5900 PMT
synthetic fused silica (SQ1) radiator
→ Cherenkov radiation
→ no lead shield necessary (?)
A. Kaukher et al.,
Proc. DIPAC 2011, TUPD21
Proc. BIW 2012, MOPG007
Libera BLD (Spark) → DC-to-8 MHz BW @ 50Ω input (DC-to-10kHz BW @ 10kΩ input)

## Slide 8: BLM Test @ ESRF
Gero Kube, DESY / MDI
single bunch injection into the main ring
scraper SCRATCH4/INT inserted
DEELS 2015 Workshop @ ALBA, 15.06.2015
→ all beam dumped into the scraper
→ one signal expected from BLM
scraper SCRATCH4/INT out, …/EXT inserted
→ signals from:
injection, crashing, passing beam
→ three signals expected from BLM
origin of measured loss signals well understood

## Slide 9: BLM Test @ ESRF
Gero Kube, DESY / MDI
dark current injection into the main ring
e-gun off
injector chain as for usual operation
screen behind septum inserted
DEELS 2015 Workshop @ ALBA, 15.06.2015
→ generate losses from single electrons
all BLMs are sensitive for single electron losses
good S/N level from DESY-type BLM
→ not the same losses observed

## Slide 10: BLM Test @ ESRF
Gero Kube, DESY / MDI
filling pattern: 16 bunches, equally distributed
internal horizontal scraper moved in
DEELS 2015 Workshop @ ALBA, 15.06.2015
h = 992, fRF = 352.202 MHz
→ Trev = 2.817 μsec, tbucket = 2.84 nsec
176.03 nsec
1 turn: 2.817 μsec
16 bunches
filling pattern nicely reproduced
losses from individual bunches resolved
BLM signal
scraper position [a.u.]
ADC samples
scraper pos. 20

## Slide 11: BLM Test @ ESRF
Gero Kube, DESY / MDI
filling pattern: 1/3rd fill
internal horizontal scraper moved in
DEELS 2015 Workshop @ ALBA, 15.06.2015
scraper position [a.u.]
ADC samples
BLM signal
h = 992, fRF = 352.202 MHz
→ Trev = 2.817 μsec, tbucket = 2.84 nsec
scraper pos. 62
0.954 nsec
2.817 μsec
331 bunches
filling pattern nicely reproduced
system bandwidth → no single bunch resolution

## Slide 12: Summary and Outlook
Gero Kube, DESY / MDI
PETRA III back in operation after 1 year shut-down
test of XFEL-BLMs at ESRF
good S/N level
→ internal/external power supply
no lead shielding required (high radiator purity)
→ at least at location of test setup
discrepancy between loss rate from BLMs and lifetime
→ measured loss rate (scraper inserted) much smaller than expected from lifetime
- small signals not detected @ 50Ω impedance, i.e. high input impedance advantageous
- influence of loss geometry (?)
beam parameters reached
DEELS 2015 Workshop @ ALBA, 15.06.2015
BLMs will be a critical system for diffraction limited storage rings
BLM tests will continue at PETRA

## Slide 13: Specs
Gero Kube, DESY / MDI
DEELS 2015 Workshop @ ALBA, 15.06.2015
PMT specs
R7600U Series:
→ 1140,- EUR / piece
→ 933,- EUR / piece
radiator
synthetic fused silica (SQ1)
slab: 30 x 220mm
j-plasma GmbH
Jena (Germany)
→ 490,- EUR / piece