## Slide 1: Pinhole Camera Issues at Diamond
L. M. Bobb, A. Morgan and G. Rehm

## Slide 2: Nov 2013: Initial Design of Pinhole Assembly After a Few Years of Beam
Aluminium holder
Stainless steel screws
Tungsten blades
“Stuff” that shouldn’t be here!

## Slide 3: Nov 2013: Proposed Mechanism
X-rays
Pinhole is in AIR
Ozone
 Oxidation of metals
X-rays from 1.4 Tesla bending magnet

## Slide 4: Nov 2013: Design Modification
Aluminium holder
Silver plated screws
Choose metals less prone to oxidation
Install air blower to prevent Ozone build-up
Add attenuation to “control” rate of growth
Gold plated Tungsten blades

## Slide 5: Nov 2014: Gold Plated Tungsten Blades (1)
50 µm
25 µm
Upstream view of Pinhole 2
Fluffy white deposits
= Aluminium Oxide
Gold “nuggets” clogging pinhole
Tungsten Oxide (blue) forms on bare Tungsten areas
400 µm

## Slide 6: Nov 2014: Gold Plated Tungsten Blades (2)
400 µm blade surface
25µm surface

## Slide 7: Nov 2014: Design Modifications
Go back to bare Tungsten blades
Remove air blower
Keep attenuation to “control” rate of growth
Develop Pinhole 3 system specifically for R&D
Potential Improvements to be investigated:
Alternative materials used for the blades:
  Solid gold or platinum (Zero oxidation but expensive)
  Machined Molybdenum so that shims aren’t required
Nitrogen purging through pinhole assembly
NB: We do not want to put the pinhole assembly in vacuum to ensure easy access and maintenance.

## Slide 8: Comparison of X-ray ports
 | DLS | ALBA | ESRF
Energy [GeV] | 3 | 3 | 6
Stored current [mA] | 300 | 100 | 200
Bend field [T] | 1.4 | 1.42 | 0.5
H angle [mrad] | 0.26 | 1 | 1
Al window [mm] | 1 | 1 | 2
Power after window [W] | 1.7 | 2.1 | 15.1
Absorber | 9.5mm Al | 0.4mm Cu | 1mm Cu
Power after absorber [mW] | 130 | 120 | 830
Peak energy after absorber [keV] | 38 | 42 | 58
Power absorbed in 0.5m air [mW] | 1.8 | 1.6 | 8.8
Thanks to G. Rehm (DLS), U. Iriso (ALBA) and F. Ewald (ESRF)

## Slide 9: Ozone Damage at ESRF
Fan near the pinhole stage failed.
The fan was supposed to blow the ozone away
New X-ray window with CVD diamond.
Weeks later…
Thanks to ESRF for these images and discussions

## Slide 10: Ozone Growth at ALBA
Main region of growth is on Copper attenuator
“Open air” mount for Tungsten blades in comparison to DLS pinhole assembly
Thanks to ALBA for the tour

## Slide 11: Thank you for your attention!